Validatable

Continuous Band Sealer

D 542 AVTSV

Equipped with:

Audion Touch Techware Version BS1.01_01

Technical

Specification Sheet

Rev.00

| General specifications |  |
| --- | --- |
| Type of machine | Validatable vertical continuous band sealer |
| Model | D 542 AVTSV |
| Machine dimensions WxDxH | 1396 x 800 x 1588 mm (with conveyor belt) |
| Weight | 100 kg |
| Seal width | 10 mm |
| Seal system | Knurled seal |
| Bag material | Sealable pouches (PE, PP, laminate, paper bags, etc.) Preformed sterile barrier system conforming ISO 11607-1, EN 868-5, EN868-9 |
| Material thickness | 60 – 200 micron (depends on the type of material) |
| Bag length | 45 – 860 mm (vertical sealing orientation with conveyor) |
| Max. carrying weight | 10 kg |
| Control system | Audion Touch Techware Equipped with Audion Temperature Controller (ATC) + 7 inch touchscreen (ADC) |
| Data transfer | USB memory stick |
| Covering | Stainless steel |
| Sealing height adjustable range (manual adjustment) | 757 - 1170 mm (vertical orientation) 917 - 1330 mm (horizontal orientation) |
| Conveyor belt height adjustable range | 388 – 838 mm (in steps of 50 mm) |
| Conveyor belt width | 250 mm |
| Power supply | 230 V – 1ph – 50Hz |
| Power consumption | 1100 W |
| Noise level | < 85 dB(A) |
| IP rating | IP20 |
| Applicable standards | 2006/42/EC, 2014/30/EU, 2011/65/EU+2015/863, EN-ISO 12100:2010, EN-ISO 13732-1:2008, EN-ISO 13850:2015, EN-ISO 14118:2018, EN-IEC 60204-1:2018, EN-ISO 11607-2:2020, ISO/TS 16775:2021, EN 868-5:2018, DIN 58953-7:2020 |

| Technical specifications |  |  |  |
| --- | --- | --- | --- |
| Recipe | Storage | 50 recipes |  |
|  | Recipe backup (.bin file through USB stick) | Copy recipe to USB Load recipe from USB |  |
| Process parameters | Seal temperature Setting range Tolerance range | 80 – 200 °C (max. 250 °C on request) 3 – 20 °C (default 5 °C) |  |
|  | Seal force Setting range Tolerance range | 5 – 80N (default 30 N) +/- 0 – 50 N (default 10 N) |  |
|  | Seal speed Setting range Tolerance range | 1,0 – 10,0 m/min. +/- 0,0 – 5,0 m/min. (default 0,5m/min.) |  |
| Adjustable parameters | Fine tuning for calibration | Seal temperature Seal force |  |
| Validation functions | Parameter monitoring (alarm stop) | Seal temperature Seal force Seal speed Seal length |  |
|  | Data logging / batch record (.csv file through USB stick) | Machine ID User name Recipe name Batch ID Parameter settings Parameter tolerances | Alarm limits Date/time Seal ID Parameters results Pass/fail Alarm code |
| System & functions | User administration (option) | User name User levels | Login password Logout time setting |
|  | User levels | Operator Supervisor Administrator | Service |
|  | Language | English German Netherlands | French Italian |
|  | Alarm history | Date/time Alarm code/descriptions |  |
|  | Counter | Bag counter (resettable) Machine counter (non-resettable) |  |

Data logging file

Machine dimensions

* The diagram includes the optional conveyor belt

| Bag size range |  |
| --- | --- |
|  | In vertical sealing orientation with conveyor |

| Option | Descriptions |
| --- | --- |
| User administration | Access control with login ID and password |
| Label printer (ETO) | For printing tracing information on labels and apply on packaging Label printer : CAB MACH2 |
| Barcode scanner (ETO) | Quick recipe change by scanning a barcode Scanner : Honeywell Voyager 1400g |
| Communication with external conveyor belt (ETO) | Output signal to external conveyor to synchronize with sealing belt |

| Accessories / Services | Codes | Descriptions |
| --- | --- | --- |
| Parameter verification | ATM-R | Seal temperature verification set |
|  | AFMR | Seal force verification set |
|  | ASM | Seal speed verification set |
| Calibration |  | Audion factory calibration |
| IQ/OQ Check | IQ/OQ VAL | 1 x IQ check for 1 machine > Operation training (max. 4 persons) > Maintenance training (max. 4 persons) OQ check for 1 type of bag > Peel test for defining seal temperature > Peel test with defined seal temperature > Dye penetration test > Seal check |
|  | OQC | Extra OQ check for each extra bag |
| Seal integrity test | SIT | Seal integrity test service > Peel test > Dye penetration test > Seal check |
|  | APT 100 | Peel tester |
|  | ASC SHEET | Seal check sheet |
|  | ASC ROLL | Seal check sheet in roll |
|  | ASC INK-B | Blue ink for dye penetration test |

| Measuring instruments for parameter verification |  |
| --- | --- |
| Seal temperature ATM-R | - Measurement range: -60°C to 350°C - Thermocouple: NiCr-Ni Type K - Diameter: 6 mm - Resolution: 0,1°C (up to 199,9°C) / 1 °C (above 200°C) - Accuracy display unit: 0.3% of reading + 1°C - Measuring rate: 2,5 per second - Minimum and maximum values readable - Operating conditions: +5 ... 50 °C, 30 – 70% RH (non-condensing) - Power supply: 9V battery - Battery life: +/- 100 hours - Dimensions: 63 (W) x 186mm (L) x 40mm (D) - Weight: 350 grams |
| Seal force AFMR | - Measurement range: max. 2000 N - Resolution: 0,1 N - Accuracy class: 0,5 % - Operating temperature range: -10 ... +50 °C - Max. Overload: 120% - Measuring rate: 50 per second - Power supply: 3x AA batteries - Battery life: +/- 40 hours - Environmental protection: IP 54 - Dimensions: 82mm (W) x 162mm (L) x 54mm (D) - Weight: 670 grams (excluding mounting tools) |
| Seal speed ASM | - Display: 5-digits, 5-alphanumeric LCD - Range (Optical): 5 to 200,000 RPM (subject to ambient light intensity) - Range (Contact): 0.5 to 20,000 RPM (also reads RPS and RPH) - Accuracy (Optical): +/-0.01% of reading - Accuracy (Contact): +/-0.05% of reading - Totalizer: 1-999,990 (events or length) - Resolution: 0.001 to 10 RPM (range dependent) - Operating Range: 2 in. to 25 ft.*, ±70° from perpendicular - Memory: Maximum, Minimum, and Last - Power: 2 AA 1.5 V dc batteries (30 hours) - Environmental: Temp: 5° to 40 °C, RH: 80% RH up to 31 °C - Dimensions (HxWxD): 17.58 x 6.10 x 4.06 cm - Weight: 210 g |